## Slide 1
ECSA-10
Early Results on GNSS Receiver Antenna Calibration System Development
Antonio Tupek 1,*, Mladen Zrinjski 1, Marko Švaco 2, Đuro Barković 1
10th International Electronic Conference on Sensors and Applications
15–30 November 2023
1 University of Zagreb, Faculty of Geodesy, Croatia
2 University of Zagreb, Faculty of Mechanical Engineering and Naval Architecture, Croatia
* Correspondence: antonio.tupek@geof.unizg.hr

## Slide 2
ECSA-10 | 15 – 30 November 2023
2
Introduction
  GNSS receiver antenna calibration
  Definitions
  Antenna Calibration Methods
Antenna Calibration System @ LMMT
  System components
  Data processing and PCC estimation
Preliminary Results
Conclusion & Outlook
PRESENTATION OUTLINE

## Slide 3: INTRODUCTION (1)
ARP
PC
Correction!
Global Navigation Satellite Systems (GNSS) positioning – 3D trilateration
GNSS satellite signals – L-band (1 – 2 GHz)
Geodetic GNSS positioning – carrier phase observations
GNSS receiver antenna electrical phase centre – PC
Antenna Reference Point – ARP
3
ECSA-10 | 15 – 30 November 2023

## Slide 4
Phase Center Correction – PCC
  Phase Center Offset – PCO
  Phase Center Variation – PCV
GNSS Antenna Calibration
Individual and type-mean PCC models
ANTEX (ANTenna EXchange)
IGS → igs20.atx
4
ECSA-10 | 15 – 30 November 2023
INTRODUCTION (2)
Definition and geometrical interpretation of the adopted GNSS receiver antenna phase center correction (PCC) model; ARP — Antenna Reference Point; MPC — Mean Phase Center; APC — Actual Phase Center; PCO — Phase Center Offset; PCV — Phase Center Variation.

## Slide 5
5
CALIBRATION METHODS
RELATIVE FIELD CALIBRATION
Rothacher et al. (1995), Mader i MacKay (1996), Mader (1999)
ABSOLUTE FIELD CALIBRATION
Wübbena et al. (1997), Wübbena et al. (2000)
CHAMBER CALIBRATION
Zeimetz (2010)
ECSA-10 | 15 – 30 November 2023
INTRODUCTION (3)

## Slide 6: ANTENNA CALIBRATION @ LMMT (1)
6
ECSA-10 | 15 – 30 November 2023
Calibration System Hardware Components:
  Industrial robot – Mitsubishi MELFA RV-4FLM-Q
  GNSS antennas – REF and AUT
  GNSS receivers – REF and AUT
  Personal computer
  Network switch, required wiring, ...
Software:
  Antenna calibration module (ACM)
  Time Synchronization module (TISY)
  PCC estimation module

## Slide 7: ANTENNA CALIBRATION @ LMMT (2)
7
ECSA-10 | 15 – 30 November 2023
REF
Robot controller
MELFA RV-4FLM-Q

## Slide 8: ANTENNA CALIBRATION @ LMMT (3)
8
ECSA-10 | 15 – 30 November 2023
Calibration setup:
  5 m baseline
  2088 different antenna orientations (rotation and tilting)
  Robot stationary for 2.5 s at every antenna orientation
  10 Hz GNSS observations
  Full antenna calibration ≈ 2 h
  Antenna orientations timestamped with GPS time (TISY)
Satellite azimuth and zenith angle plot (sky-plot) during antenna calibration at LMMT; (a) topocentric frame (TF); (b) antenna frame (AF).

## Slide 9: ANTENNA CALIBRATION @ LMMT (4)
9
ECSA-10 | 15 – 30 November 2023
Triple-difference (TD) approach:
  Single (station) differences (SD)
  Double (satellite) differences (DD)
  Triple (time) differences (TD)
Spherical Harmonics (SH) parametrisation:
ANTEX output

## Slide 10: PRELIMINARY RESULTS (1)
10
ECSA-10 | 15 – 30 November 2023
4 calibration campaigns (April – June 2023)
Trimble Zephyr 2 Geodetic (TRM57971.00 NONE, S/N:30739001)
Global Positioning System (GPS) L1 frequency (G01)
Validation with Geo++ GmbH

## Slide 11: PRELIMINARY RESULTS (2)
11
ECSA-10 | 15 – 30 November 2023
Calibration Institution | Phase Center Offset (PCO) |  | 
 | North | East | Up
Geo++ GmbH | 0.79 | 0.32 | 66.96
LMMT | 1.24 | 0.11 | 67.24
PCO vector components of antenna TRM57971.00 NONE (S/N: 30739001) of GPS L1 (G01) frequency per calibration institution. All values are in millimeters (mm).
(a) Geo++ GmbH and (b) LMMT phase center variations (PCVs) of antenna TRM57971.00 NONE (S/N: 30739001) for GPS L1 (G01) frequency.

## Slide 12: PRELIMINARY RESULTS (3)
12
ECSA-10 | 15 – 30 November 2023
Full antenna hemisphere (0° elevation cut-off) |  |  |  |  | Reduced antenna hemisphere (10° elevation cut-off) |  |  |  | 
Min. | Max. | RMS | Range | IQR | Min. | Max. | RMS | Range | IQR
–2.92 | 0.72 | 0.52 | 3.64 | 0.44 | –1.64 | 0.72 | 0.34 | 2.36 | 0.41
Quantitative measures of the difference between Geo++ GmbH and LMMT individual absolute calibrations. All values are in millimetres (mm).
Geo++ GmbH and LMMT PCC differences (ΔPCC) of antenna TRM57971.00 NONE (S/N: 30739001) for GPS L1 (G01) frequency after transformation to common PCO and datum.

## Slide 13
13
CONCLUSION & OUTLOOK
Antenna Calibration System is operational and provides meaningful receiver antenna calibration results for GPS L1 frequency
Submillimeter accuracy achieved
  Validation with Geo++ GmbH to within 0.52 mm
Robot control software and time synchronization – successful
Outlook: calibration sequence optimization, further system development with the end goal of providing multi-GNSS calibrations
ECSA-10 | 15 – 30 November 2023

## Slide 14: REFERENCES
14
ECSA-10 | 15 – 30 November 2023
Mader, G.L. GPS Antenna Calibration at the National Geodetic Survey. GPS Solut. 1999, 3, 50–58. https://doi.org/10.1007/PL00012780.
Rothacher, M., Schaer, S., Mervart, L., Beutler, G. (1995): Determination of Antenna Phase Center Variations using GPS Data, in Proceedings of the IGS Workshop – Special Topics and New Directions, May 15–18, Potsdam, Germany, 205–220.
Wübbena, G.; Schmitz, M.; Menge, F.; Seeber, G.; Völksen, C. A New Approach for Field Calibration of Absolute GPS Antenna Phase Center Variations. Navigation 1997, 44, 247–255. https://doi.org/10.1002/j.2161-4296.1997.tb02346.x.
Wübbena, G.; Schmitz, M.; Menge, F.; Böder, V.; Seeber, G. Automated Absolute Field Calibration of GPS Antennas in Real-Time. In Proceedings of the 13th International Technical Meeting of the Satellite Division of The Institute of Navigation (ION GPS 2000), Salt Lake City, Utah, USA, 19–22 September 2000; The Institute of Navigation: Manassas, VA, USA, 2000; pp. 2512–2522.
Zeimetz, P. Zur Entwicklung Und Bewertung Der Absoluten GNSS-Antennenkalibrierung Im HF-Labor. Ph.D. Thesis, University Bonn, Bonn, Germany, 2010.

## Slide 15
Thank you for your attention!
Antonio Tupek, MSc
Research and teaching assistant
Faculty of Geodesy, University of Zagreb
Zagreb, Croatia
antonio.tupek@geof.unizg.hr